SCOPE 3 ЛК

Metrolab OOD

laboratory Complex Interlab

UIC: 206477195

Management Address: 1359 Sofia, Lyulin, Bl. 408, Ent. A, Floor 2, Apt. 11 
Laboratory Address: 1574 Sofia, 13 Prof. Tsvetan Lazarov Blvd.

To perform calibrating of:

| Type of the the scope: Fixed |
| --- |
| № | Measuring instrument* | Measured value, unit of measure | Measurement Range | Measurement uncertainty | Calibration method |
| 1 | 2 | З | 4 | 5 | 6 |
| 1. | Non-automatic weighing instruments | mass, kg |  |  | LKI 08- 05:2020 |
|  | Accuracy class I |  | Up to 0,5 kg | 1,8.10-4 g |  |
|  | Accuracy class II |  | From 0,5 kg to 60 kg | From 1,8.10-4 g to 2 g |  |
|  | Accuracy class III |  | from 60 kg to 150 kg | from 2 g to 5 g |  |
| 2. | Devices for measuring air flow velocity | Velocity, m/s | from 0,137 m/s to 2,000 m/s | from 0,060 m/s to 0,095 m/s | LKI 10- 05:2020 |
|  |  |  | from 2,000 m/s to 30,17 m/s | from 0,095 m/s to 0,51 m/s |  |
| 3. | Liquid-in-glass thermometers | Temperature, °С | From minus 50°С to 20°С | from 0,3°С to 0,09°С | LKI 13- 11:2020 |
|  |  |  | from 20°С to 200°С | from 0,09 °С to 0,3 °С |  |
|  |  |  | from 200°С to 600°С | from 0,3°С to 1,2 °С |  |
| 4. | Digital thermometers Manometric and bimetallic thermometers | Temperature, °С | From minus 50°С to 20°С | from 0,3°С to 0,09°С | LKI 14- 11:2020 |
|  |  |  | from 20°С to 200°С | from 0,09°С to 0,3 °С |  |
|  |  |  | from 200°С to 1200 °С | from 0,3°С to 2,5°С |  |
| 5. | Thermocouples | Temperature, °С | From minus 50°С to 1200 °С | from 1,2°С to 2,4 °С | LKI 11- 05:2020 |
| 6. | Ultraviolet and visible spectrophotometers | Spectral transmittance ТΛ, %, geometry % to air | from 8 % to 93% with λ from 257 nm to 690 nm | from 0,1 % to 1% | LKI 07- 04:2020 |
|  |  | Spectral optical density, Dλ, geometry % to air | from 0,03 to 1.1 with λ from 257 nm to 690 nm | от 5.10-3 to 8.10-3 | LKI 07- 04:2020 Calculation method |
|  |  | Wavelength, λ, nm | from 241 nm to 638 nm | from 0,3 nm to 1 nm | LKI 07- 04:2020 |
| 7. | pH meters | Hydrogen indicator | from 2 to 12 | from 1,5.10-2 to 3,10-2 | LKI 09- 05:2020 |
| 8. | Conductometers | Specific concuctivity mS/cm | from 1.3.10-3 mS/cm to 111,3.10-3 mS/cm | from 0,4.10-5 mS/cm to 0,8.10-5 mS/cm | LKI 01- 09:2020 |
| 9. | Calorimeters | Heat of combustion, J/g | from 26400 J/g to 26600 J/g | 6 J/g | LKI 04- 09:2020 |
| 10. | Devices for measuring relative air humidity | Relative humidity %rh | from 20 %rh to 85 %rh | from 1,2%rh to 1.4%rh | LKI 12- 09:2020 |

References:

- LKI 01-09:2020 Methodology for calibration of conductometers.

- LKI 04-09:2020 Methodology for calibration of calorimeters

- LKI 07-04:2020 Methodology for calibration of spectrophotometers

- LKI 08-05:2020 Methodology for calibration of non-automatic weighing instruments

- LKI 09-05:2020 Methodology for calibration of pH meters

- LKI 10-05:2020 Methodology for calibration of devices for measuring air flow velocity

- LKI 11-05:2020 Methodology for calibration of thermocouples

- LKI 12-09:2020 Methodology for calibration of devices for measuring relative air humidity

- LKI 13-11:2020 Methodology for calibration of liquid-in-glass thermometers

- LKI 14-11:2020 Methodology for calibration of digital thermometers, manometric and bimetallic thermometers

Notes:

- The calibration of the measurement devices in items 2, 3, 4 and 10 was carried out in the laboratory.

- The calibration of the measurement devices in items 5, 7 and 8 was carried out in the laboratory or at the client’s site.

- The calibration of the measurement devices in 1, 6 and 9 was carried out at the client’s site