SCOPE 22 ЛК

DELTA INSTRUMENTS LTD.

CALIBRATION LABORATORY OF MEASUREMENT EQUIPMENT

FOR ELECTRICAL QUANTITIES

Management and laboratory address:

1616 Sofia, Boyana, Vitosha district, 20, 732 Str.

To perform calibrating of:

| Type of the scope: Fixed |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| № | Measuring Instrument | Measure and, Measure ment Unit | Measurement Range | Measurement Uncertainty | Calibration Method |
| 1 | 2 | 3 | 4 | 5 | 6 |
| 1. | DC Voltmeters | DC Voltage, V | 20 mV to 100 mV 100 mV to 1 V 1 V to 10 V 10 V to 100 V 100 V to 1000 V | 1,6.10-2 mV to 3,0.10-2 mV 3,0.10-2 mV to 1,3.10-4 V 1,2.10-4 V to 2,5.10-3 V 2,5.10-3 V to 1,3.10-2 V 1,3.10-2 V to 0,2 V | CP-702-02-06/2019 |
| 2. | AC Voltmeters | AC Voltage, (50 Hz) V | 0,2 V to 1 V 1 V to 10 V 10 V to 100 V 100 V to 1000 V | 4.10-4 V to 1,5.10-3 V 1,5.10-3 V tо 1,5.10-2 V 1,5.10-2 V to 0,16 V 0,16 V to 1,9 V | CP-702-02-06/2019 |
| 3. | DC Ammeters testers | DC current, А | 0,2 mA to 2 mA 2 mA to 20 mA 20 mA to 200 mA 0,2 A to 2 А 2 A to 10 А | 3,2.10-4 to 2,7.10-3 mA 2,7.10-3 mA to 2,7.10-2 mA 2,7.10-2 mA to 0,24 mA 2,4.10-4 A to 2,7.10-3 A 2,7.10-3 A to 0,014 A | CP-702-02-06/2019 |
| 4. | AC Ammeters testers | AC current, (50Hz) А | 2 mA to 20 mA 20 mA to 200 mA 0,2 A to 2 А 2 A to 10 А | 1,3.10-2 mA to 0,071 mA 0,071 mA to 0,6 mA 6,0.10-4 A to 7.10-3 A 7.10-3 A to 0,052 A | CP-702-02-06/2019 |
| 5. | DC resistance meters (Ohmmeters) | Electrical Resistance, Ω | 0,5 Ω to 10 Ω 10 Ω to 100 Ω 100 Ω to 1000 Ω 1 kΩ to 10 kΩ от 10 kΩ to 100 kΩ | 1.10-4 x R + 6 mΩ 3.10-4 x R + 4 mΩ 3.10-4 x R + 4 mΩ 3.10-4 x R 3.10-4 x R | CP-702-02-03/2018 |
| 6. | Earth testers (Instruments for measuring the parameters of electrical installations by earth resistance) | Electrical Resistance, Ω | 0,5 Ω tо 1 Ω 1 Ω tо 20 Ω 20 Ω tо 200 Ω 200 Ω tо 2000 Ω | 10 mΩ 10 mΩ 5.10-4 x R 5.10-4 x R | CP-702-02-05/2018 |
| 7. | Loop impedance testers (Instruments for measuring the parameters of electrical installations by loop resistance) | Electrical Resistance, Ω | 0,1 Ω tо 19,99 Ω 20 Ω tо 200 Ω | 2,5.10-3 x R + 0,05 Ω 1.10-3 x R + 0,09 Ω | CP-702-02-04/2018 |
| 8. | Residual-current device (RCD testers) by contact voltage (Instruments for measuring the parameters of electrical installations by contact voltage) | AC Voltage, (50Hz) V | 0,1 V to 9,99V 10 V to 49,9 V 50 V to 90 V | 5.10-2 x U + 0,1 V 6.10-2 x U + 0,02 V 5,2.10-2 x U + 0,6 V | CP-702-02-02/2018 |
| 9. | Residual-current device (RCD testers) by tripping current (Instruments for measuring the parameters of electrical installations by tripping current) | AC current, (50Hz) А | 10 mA 30 mA 100 mA 300 mA 500 mA | 0,1 mA 0,1 mA 0,2 mA 1 mA 2 mA | CP-702-02-02/2018 |
| 10. | Insulation testers | Electrical Resistance, Ω | 1 kΩ, 10 kΩ, 100 kΩ 1 MΩ, 10 MΩ,100 MΩ 1 GΩ, 10 GΩ 100 GΩ 1 TΩ | 5.10-4 x R 15.10-3 x R 15.10-3 x R 5.10-2 x R 5.10-2 x R | CP-702-02-01/2018 |
| 11. | Instruments for measuring the parameters of electrical installations by tripping time | Time, s | 10 ms 30 ms 100 ms 200 ms | 0,8 ms 1 ms 1 ms 1 ms | CP-702-02-07/2023 |
| 12. | AC/DC Current clumps | AC current, (50Hz) А | 2 mA to 20 mA 20 mA to 200 mA 200 mA to 2 A 2 A to 10 A | 0,02 mA to 0,08 mA 0,08 mA to 0,95 mA 0,95 mA to 0,01 А 0,01 А to 0,1 А | CP-702-02-08/2023 |
|  |  | AC current, (50Hz) А | 10 А to 100 А 100 А to 500 А 500 А to 950 А | 0,1 А to 1,0 А 1,0 А to 4,7 А 4,7 А to 8,3 А |  |
|  |  | DC current, А | 10 А to 100 А 100 А to 500 А 500 А to 950 А | 0,1 А to 0,8 А 0,8 А to 4,0 А 4,0 А to 7,9 А |  |
| 13. | Frequency measuring instruments | Frequency, Hz | 2 Hz to 200 kHz | 0,078 Hz to 0,015 kHz | CP-702-02-09/2024 |
| 14. | Power quality analyzers | AC Voltage (50 Hz), V | 100 V to 600 V | 0,16 to 1,2 V | CP-702-02-10/2024 |
|  |  | AC Current (50 Hz) A | 2 А to 10 А 10 А to 100 А 100 А to 950 А | 0,0046 А to 0,053 А 0,053 А to 0,82 А 0,82 А to 8,4 А |  |

References:

| CP-702-02-01/2018 | Calibration Procedure of measurement equipment for electrical quantities in insulation resistance (megohmmeters, meggers, insulation testers). |
| --- | --- |
| CP-702-02-02/2018 | Calibration Procedure of measurement equipment for electrical quantities in trip current of RCDs and contact voltage. |
| CP-702-02-03/2018 | Calibration Procedure of measurement equipment for electrical quantities in resistance. |
| CP-702-02-04/2018 | Calibration Procedure of measurement equipment for electrical quantities in loop impedance. |
| CP-702-02-05/2018 | Calibration Procedure of measurement equipment for electrical quantities in ground resistance. |
| CP-702-02-06/2019 | Calibration Procedure of measurement equipment for electrical quantities in direct current and voltage, alternating current and voltage. |
| CP-702-02-07/2023 | Calibration Procedure of measurement the parameters of the electrical installations for tripping time. |
| CP-702-02-08/2023 | Calibration Procedure of measurement current clupms. |
| CP-702-02-09/2024 | Calibration Procedure of frequency measuring instruments |
| CP-702-02-10/2024 | Calibration Procedure of power quality analyzers |

Note 1: Calibration of all measuring instruments is carried out at the laboratory premises.

Note 2: Calibrated measuring instruments may be stand-alone or part of combined devices (including multimeters).

Note 3: Insulation resistance measuring devices can be megohmmeters, meggers and insulation testers.

Note 4: Calibrated measuring instruments may be analogue or digital.