SCOPE 23 ЛК

Sofia, 28.07.2025

EMSYST-6 LTD.

CALIBRATION LABORATORY EMSYST

Management and Laboratory address:

Bulgaria, 1784 Sofia, 133 Tsarigradsko Shosse Blvd, BIC IZOT, Office 304

To perform calibrating of:

| Type of the scope: Fixed |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| № | Measuring Instrument | Measure and, Measure ment Unit | Measurement Range | Measurement Uncertainty | Calibration Method |
| 1 | 2 | 3 | 4 | 5 | 6 |
| 1. | Standard Electricity Meters, Electronic, Single-Phase and Three-Phase for Active Energy at frequency 50 Hz | Electrical Energy, Active, kWh | Per phase From 1,25 Ws to 21,6.106 Ws Voltage (U): From 50 V to 300 V Current (I): From 0,05 A to 120 A Power Factor: From 1 to 0,5 ind, or from 1 to 0,8 cap Time interval from 1 s to 600 s | 0,020 % at cos φ = 1 U ≤ 230 V | WI 7.6.1-1 № E-MK-01/20 |
|  |  |  |  | 0,025 % at cos φ = 1 U>230 V and at cos φ =0,5 ind/ cos φ =0,8 cap U ≤ 230 V I ≤ 12 A |  |
|  |  |  |  | 0,030 % at cos φ = 0,5 ind/ cos φ = 0,8 cap I > 12 A |  |
| 2. | Standard Electricity Meters, Electronic, Single-Phase and Three-Phase for Reactive Energy at frequency 50Hz | Electrical Energy, Reactive, kvarh | Per phase from 0,625 vars to 21,6.106 vars Voltage (U) from 50 V to 300 V Current (I) from 0,05 A to 120 A Power Factor from 1 to 0,25, ind or cap Time interval From 1s to 600s | 0,025 % at sin φ =1 U ≤ 230 V | WI 7.6.1-1 № E-MK-01/20 |
|  |  |  |  | 0,030 % at sin φ =1 U > 230 V, and at sin φ =0,25 ind/cap U ≤ 230 V I ≤ 12 A |  |
|  |  |  |  | 0,035 % at sin φ =0,25 ind/cap I > 12 A |  |
| 3. | Test Benches with an Electronic Standard Electricity Meter for Metrological Verification of Electricity Meters, Single-Phase and Three-Phase, for Active and Reactive Energy at frequency 50Hz | Electrical Energy, Active, kWh, and Reactive, kvarh | For active energy, per phase from 1,25 Ws to 21,6.106 Ws Voltage (U) from 50 V to 300 V Current (I) from 0,05 A to 120 A Power Factor From 1 to 0,5 ind, or from 1 to 0,8 cap Time interval from 1 s to 600 s | 0,020 % at cos φ =1 U ≤ 230 V | WI 7.6.1-4 № EУ-МК-04/25 |
|  |  |  |  | 0,025 % at cos φ =1 U > 230 V and at cos φ =0,5 ind/ cos φ =0,8 cap U ≤ 230 V I ≤ 12 A |  |
|  |  |  |  | 0,030 % at cos φ =0,5 ind/ cos φ =0,8 cap I > 12 A |  |
|  |  |  | For reactive energy, per phase From 0,625 vars to 21,6.106 vars Voltage (U) from 50 V to 300 V Current (I) from 0,05 A to 120 A Power factor from 1 to 0,25 ind or cap Time interval from 1 s to 600 s | 0,025 % at sin φ = 1 U ≤ 230 V |  |
|  |  |  |  | 0,030 % at sin φ = 1 U > 230 V and at sin φ = 0,25 ind/cap U ≤ 230 V I≤ 12 A |  |
|  |  |  |  | 0.035 % at sin φ =0,25 ind/cap I >12 A |  |
| 4. | Flow Meters and Portable Flow Meter Stations, Calibrated with Operating Fluid Water in the range from 0,006 m3/h to 70,00 m3/h | Volume, m3 | From 0,001 m3 to 0,3 m3 (at the range from 0,006 m3/h to 30,0 m3/h) (at the range from 30,0 m3/h to 70,0 m3/h) | 0,10 % 0,20% | WI 7.6.1-2 № P-MK-01/20 |

References:

1. WI 7.6.1-1 № E-MK-01/20 Calibration Methodology for Standard Electronic Electricity Meters, validated on 17.07.2020;

2. WI 7.6.1-4 № ЕУ-МК-04/25 Calibration Methodology for Test Benches with an Electronic Standard Electricity Meter for Metrological Verification of single-phase and three-phase electricity meters for active and reactive energy, validated on 22.05.2025;

3. WI 7.6.1-2 № P-MK-01/20 Calibration Methodology for Flow Meters and Portable Flow Meter Stations, validated on 03.09.2020.

Note:

The calibrations of measurement instruments for positions 1, 2 and 3 shall be carried out in the Laboratory premises, and on the customer’s site.

The calibrations of measurement instruments for position 4 shall be carried out only in the Laboratory premises.