SCOPE 27 ЛК

AERO TECHNIC BG LTD.
 IN LAB AERONAUTIC CALIBRATION LABORATORY

Management and laboratory address:

1839, Sofia, Vrazhdebna, 8-th Str,  Hangar Heli ER SAU

To perform calibrating of:

| Type of the scope: fixed |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| № | Measuring Instrument | Measured value and, Measure ment Unit | Measurement Range | Measurement Uncertainty | Calibration Method |
| 1 | 2 | 3 | 4 | 5 | 6 |
| Electrical Quantities |  |  |  |  |  |
| 1. | Voltmeters for DC voltage (1),(2) | DC Voltage (DCU), V | from 0,02 V to 0,2 V from 0,2 V to 2 V from 2 V to 20 V from 20 V to 200 V from 200 V to 1000 V | from 5,9E-05 V to 6,4E-05 V from 6,4E-05 V to 6,1E-04 V from 6,1E-04 V to 6,1E-03 V from 6,1E-03 V to 6,3E-02 V from 6,3E-02 V to 5,9E-01 V | РПК-7.2.1-03:2024 |
| 2. | Voltmeters for AC voltage (1),(2) | AC Voltage (ACU), V | At 50 Hz frequency: from 0,02 V to 0,2 V from 0,2 V to 2 V from 2 V to 20 V from 20 V to 200 V from 200 V to 1000 V At 1 kHz frequency from 0,02 V to 0,2 V from 0,2 V to 2 V from 2 V to 20 V from 20 V to 200 V from 200 V to 1000 V | from 1,2E-04 V to 5,4E-04 V from 5,4E-04 V to 5,3E-03 V from 5,3E-03 V to 5,1E-02 V from 5,1E-02 V to 4,2E-01 V from 4,2E-01 V to 1,4 V from 1,3E-04 V to 5,4E-04 V from 5,4E-04 V to 5,4E-03 V from 5,4E-03 V to 5,1E-02 V from 5,1E-02 V to 4,2E-01 V from 4,2E-01 V to 1,8 V | РПК-7.2.1-03:2024 |
| 3. | Ampere meters for DC current | Direct Current DCI, A | from 20 µA to 200 µA from 0,2 mA to 2 mA from 2 mA to 20 mA from 20 mA to 200 mA from 0,2 A to 2 A | from 9,5Е-02 µA to 2,7Е-01 µA from 2,7Е-04 mA to 2,7Е-03 mA from 2,7Е-03 mA to 2,3Е-02 mA from 2,3Е-02 mA to 5,2Е-01 mA from 5,2Е-04 A to 3,3Е-03 A | РПК-7.2.1-03:2024 |
| 4. | Ampere meters for AC current (1), (2) | Alternating current ACI, A | At 50 Hz frequency from 100 µA to 200 µA from 0,2 mA to 2 mA from 2 mA to 20 mA from 20 mA to 200 mA from 0,2 A to 2 A At 1 кHz frequency from 100 µA to 200 µA from 0,2 mA to 2 mA from 2 mA to 20 mA from 20 mA to 200 mA from 0,2 A to 2 A | from 3,2Е-01 µA to 8,6Е-01 µA from 8,6Е-04 mA to 5,6Е-03 mA from 5,6Е-03 mA to 4,2Е-02 mA from 4,2Е-02 mA to 4,2Е-01 mA from 4,2Е-04 A to 1,5Е-02 A from 3,3Е-01 µA to 8,6Е-01 µA from 8,6Е-04 mA to 6,0Е-03 mA from 6,0Е-03 mA to 4,2Е-02 mA from 4,2Е-02 mA to 4,2Е-01 mA from 4,2Е-04 A to 1,6Е-02 A | РПК-7.2.1-03:2024 |
| 5. | Digital Ohmmeters, (1) | Active Resistance R, Ω | from 0,1 Ω to 1 Ω from 1 Ω to 10 Ω from 10 Ω to 100 Ω from 100 Ω to 1000 Ω from 1 kΩ to 10 kΩ from 10 kΩ to 100 kΩ from 0,1 MΩ to 1 MΩ from 1 MΩ to 10 MΩ from 10 MΩ to 100 MΩ from 100 MΩ to 1000 MΩ | from 2,9Е-05 Ω to 2,9Е-04 Ω from 2,9Е-04 Ω to 2,3Е-03 Ω from 2,3Е-03 Ω to 2,3Е-02 Ω from 2,3Е-02 Ω to 2,3Е-01 Ω from 2,3Е-04 kΩ to 2,3Е-03 kΩ from 2,3Е-03 kΩ to 2,3Е-02 kΩ from 2,3Е-05 MΩ to 2,6Е-04 MΩ from 2,6Е-04 MΩ to 2,4Е-03 MΩ from 2,4Е-03 MΩ to 2,3Е-02 MΩ from 2,3Е-02 MΩ to 2,4Е-01 MΩ | РПК-7.2.1-07:2024 |
| 6. | Digital instruments measuring capacitance (1) | Capacitance C, F | At 1 kHz frequency from 1 nF to 10 nF from 10 nF to 100 nF from 0,1 µF to 1 µF from 1 µF to 10 µF from 10 µF to 70 µF | from 1,5Е-02 nF to 1,5Е-01 nF from 1,5Е-01 nF to 1,5 nF from 1,5Е-03 µF to 1,6Е-02 µF from 1,6Е-02 µF to 1,6Е-01 µF from 1,6Е-01 µF to 6,0Е-01 µF | РПК-7.2.1-08:2024 |
| 7. | Digital frequency meters (1) | Frequency f, Hz | from 1 Hz to 100 Hz from 0,1 kHz to 1 kHz from 1 kHz to 100 kHz from 0,1 MHz to 1 MHz from 1 MHz to 100 MHz | from 6,3Е-03 Hz to 5,0Е-02 Hz from 5,0Е-05 kHz to 1,1Е-05 kHz from 1,1Е-05 kHz to 1,6Е-04 kHz from 1,6Е-07 MHz to 1,5Е-06 MHz from 1,5Е-06 MHz to 1,6Е-04 MHz | РПК-7.2.1-09:2024 |
| Linear quantities |  |  |  |  |  |
| 8. | Caliper | Length l, m | 0,5 mm to 300 mm | 0,03 mm | РПК-7.2.1-01:2024 |
| 9. | Cylindrical length gauges – GO/NO-GO gauges | Diameter d, m | from 0,1 mm to 30 mm | 0,002 mm | РПК-7.2.1-06:2024 |
| Torque |  |  |  |  |  |
| 10. | Torque wrenches and screwdrivers | Torque, Nm | from 0,4 Nm to 20 Nm from 20 Nm to 200 Nm | 1% 1% | РПК-7.2.1-02:2024 |
| Pressure |  |  |  |  |  |
| 11. | Electromechanical and mechanical manometers | Pressure P, Pa | up to 69 MPa with water or oil from 27 hPa to 3500 hPa with air | from 0,05% FS (4) from 0,2% FS (5) from 0,08 hPa to 0,32 hPa (6) | РПК-7.2.1-04:2024 (4) – standard procedure, (5) – basic procedure, and (6) – precision procedure |
| Mass |  |  |  |  |  |
| 12. | Non-automatic scales, class III and IV (3) | Mass m, kg | From 10 g to 30 000 g | from 0,58 g to 2,8 g | РПК-7.2.1-05:2024 |

Note:

(1) – including as part of combined instruments (multimeters);

(2) – digital and analog;

(3) – calibrations for these items are performed only at the customer's site. All other calibrations are performed only in the laboratory premises.

References:

1. РПК-7.2.1-01:2024 - Calibration of calipers (based on VDE/VDI/DGQ 2618, item 9.1 and БДС EN ISO 13385-2).

2. РПК-7.2.1-02:2024 - Calibration of dynamometric tools - wrenches and screwdrivers (based on БДС EN ISO 6789-1 and БДС EN ISO 6789-2).

3. РПК-7.2.1-03:2024 - Calibration of digital and analog instruments for measuring electrical quantities – direct current, alternating current, and voltage.

4. РПК-7.2.1-04:2024 - Calibration of electromechanical and mechanical manometers (based on EURAMET Calibration Guide № 17 Version 4.1, including (4) – standard procedure, (5) – basic procedure, and (6) – precision procedure).

5. РПК-7.2.1-05:2024 - Calibration of non-automatic scales (based on EURAMET Calibration Guide № 18 Version 4.0).

6. РПК-7.2.1-06:2024 - Calibration of cylindrical length measures – GO/NO-GO gauges (based on EURAMET cg-6 Version 3.0).

7. РПК-7.2.1-07:2024 - Calibration of digital and analog instruments for measuring electrical quantities – resistance.

8. РПК-7.2.1-08:2024 - Calibration of digital and analog instruments for measuring electrical quantities – capacitance.

9. РПК-7.2.1-09:2024 - Calibration of digital and analog instruments for measuring electrical quantities – frequency.