SCOPE 14 ЛК

MICROSYST LTD,

Calibration Laboratory

Management  address: 4002 Plovdiv, 57 Trakia Str.

Laboratory address: 4002 Plovdiv, 4 Murgash Str.

To perform calibrating of:

| Range type: fixed |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| № | Measuring instrument | Measured value, unit of measure | Measurement Range | Uncertainty of measurement | Calibration method |
| 1 | 2 | 3 | 4 | 5 | 6 |
| 1 | Calibrators and voltmeters for DC voltage | Electric voltage, volt (V) | 0 V to 0,1 V | 3,5.10-6 V to 8,5.10-6 V | МК 7.02 01-01:2020 direct and comparative method |
|  |  |  | 0,1 V to 10 V | 8,5.10-6 V to 0,1.10-3 V |  |
|  |  |  | 10 V to 750 V | 0,1.10-3 V to 15.10-3 V |  |
|  | Calibrators and voltmeters for AC voltage, at industrial frequency |  | 0,01 V to 10 V | 15.10-6 V to 1,5.10-3 V |  |
|  |  |  | 10 V to 750 V | 1,5.10-3 V to 0,15 V |  |
|  | Calibrators and voltmeters for AC voltage at 1 kHz frequency |  | 0,01 V to 10 V | 33.10-6 V to 7.10-3 V |  |
|  |  |  | 10 V to 750 V | 7.10-3 V to 0,15 V |  |
| 2 | Calibrators and Ammeters for DC Current. | Electric current, ampere (А) | 0 А to 1.10-3 А | 25.10-9 А to 0,1.10-6 А | МК 7.02 01-02:2020 direct and comparative method |
|  |  |  | 1.10-3 А to 0,1 А | 0,1.10-6 А to 9.10-6 А |  |
|  |  |  | 0,1 А to 2 А | 9.10-6 А to 0,3.10-3 А |  |
|  | Calibrators and ammeters for AC Current, at industrial frequency |  | 2.10-3 A to 0,1 А | 3.10-6 А to 35.10-6 А |  |
|  |  |  | 0,1 А to 2 А | 35.10-6 А to 3.10-3 А |  |
|  | Calibrators and ammeters for AC Current, at 1kHz frequency |  | 2.10-3 A to 0,1 А | 4.10-6 А to 55.10-6 А |  |
|  |  |  | 0,1 А to 2 А | 55.10-6 А to 3.10-3 А |  |
| 3 | DC resistance standards, DC resistance calibrators and DC resistance decade boxes | Electrical resistance, ohm (Ω) | 0,01 Ω to 100 Ω | 4.10-3 Ω to 12.10-3 Ω | МК 7.02 01-03:2018 direct method |
|  |  |  | 0,1 kΩ to 100 kΩ | 12.10-6 kΩ to 3.10-3 kΩ |  |
|  |  |  | 0,1 MΩ to 100 MΩ | oт 3.10-6 MΩ to 0,2 MΩ |  |
| 4 | Ohmmeters | Electrical resistance, ohm (Ω) | 0,01 Ω to 100 kΩ | 2.10-3 Ω to 7. 10-3 kΩ | МК 7.02 01-04:2020 direct method |
|  |  |  | 1 MΩ | 0,4.10-3 MΩ |  |
|  |  |  | 10 MΩ | 2.10-3 MΩ |  |
|  |  |  | 100 MΩ | 0,24 MΩ |  |
| 5* | Converters with DCU and R input /temperature indicators/ | Temperature, degrees Celsius (°С) | For thermistors: input: 0,01 Ω to 100 kΩ output: -200°С to 850 °С | 0,06 °С | МК 7.02 01-05:2020 direct method |
|  |  |  | For thermocouples: input: -50 mV to 150 mV output: -200 °С to 1 850 °С | 0,2 °С |  |
|  | Converters with DCU input, DCI for temperature, relative humidity, pH, pressure /indicators with uniform input signal/ | Temperature, degrees Celsius (°С) | -200 °С to 1 850 °С | 0,2 °С |  |
|  |  | Relative humidity, %RH | 0 % rh to 100 % rh | 0,1 %rh |  |
|  |  | pH-hydrogen index | 0 pH to 14 pH | 0,001 pH |  |
|  |  | Pressure bar (bar) | -1 bar to 1 000 bar | 1.10-5 bar to 1.10-3 bar |  |
|  | Secondary converters with input DCU, DCI, R and ACU, ACI at industrial frequency and with output DCU, DCI | Electric voltage, volt (V) | input: 0 V DC to 750 V DC output: 0 V DC to 11 V DC | 15.10-6 V DC to 0,5.10-3 V DC |  |
|  |  | Electric current, ampere (А) | input: 0 V DC to 750 V DC output: 0 mА DC to 100 mА DC | 4. 10-3 mA DC to 15. 10-3 mА DC |  |
|  |  | Electric voltage, volt (V) | input: 0.01 V AC to 750 VAC output: 0 V DC to 11 V DC | 25.10-6 V DC to 0,5.10-3 V DC |  |
|  |  | Electric current, ampere (А) | input: 0.01 V AC to 750 V AC output:0 mА DC to 100 mА DC | 1. 10-3 mА DC to 0,2 mА DC |  |
|  |  | Electric voltage, volt (V) | input: 0 А DC to 2 А DC output: 0 V DC to 11 V DC | 15.10-6 V DC to 0,5.10-3 V DC |  |
|  |  | Electric current, ampere (А) | input: 0 А DC to 2 А DC output: 0 mА DC to 100 mА DC | 4. 10-3 mA DC to 15. 10-3 mА DC |  |
|  |  | Electric voltage, volt (V) | input: 2. 10-3 А AC to 2 А AC output: 0 V DC to 11 V DC | 25.10-6 V DC to 0,5.10-3 V DC |  |
|  |  | Electric current, ampere (А) | input: 2. 10-3 А AC to 2 А AC output: 0 mА DC to 100 mА DC | 2. 10-3 mА DC to 0,2 mА DC |  |
|  |  | Electric voltage, volt (V) | input: 0,01 Ω to 100 кΩ output: 0 V DC to 11 V DC | 18.10-6 V DC to 0,5.10-3 V DC |  |
|  |  | Electric current, ampere (А) | input: 0,01 Ω to 100 кΩ output: 0 mА DC to 100 mА DC | 0,1. 10-3 mА DC to 9. 10-3 mА DC |  |
|  | DCU, DCI and R calibrators for non-electrical values simulation | Electric voltage, volt (V) | 0 V to 11 V DC | 3,5.10-6 V DC to 0,15.10-3 V DC |  |
|  |  | Electric current, ampere (А) | 0 А to 0,1 А DC | 25. 10-9 А DC to 9. 10-6 А DC |  |
|  |  | Electric resistance, ohm (Ω) | 1 Ω to 4 .103 Ω | 5.10-3 Ω to 0,11 Ω |  |
| 6 | Digital thermometers | Temperature, degrees Celsius (°С) | -45 °C to 150 °C | 0,05 °C | МК 7.02 01-06:2018 comparative method |
|  |  |  | 150 °C to 300 °C | 0,05 °C to 0,15 °C |  |
|  |  |  | 300 °C to 650 °C | 0,15 °C to 0,7 °C |  |
| 7 | Relative humidity meters (hygrometers). | Relative humidity, (%RH) | 10 %rh to 90 %rh | 1,5 %rh to 1,8 %rh | МК 7.02 01-07:2019 comparative method |
| 8 | Thermometers (analogue and liquid) | Temperature, degrees Celsius (°С) | -45 °C to 650 °C | 0,1 °C to 0,7 °C | МК 7.02 01-08:2018 comparative method |
| 9 | Thermocouples | Temperature, degrees Celsius (°С) | -45 °C to 150 °C | 0,4 °C | МК 7.02 01-09:2018 direct method |
|  |  |  | 150 °C to 300 °C | 0,4 °C to 0,6 °C |  |
|  |  |  | 300 °C to 650 °C | 0,6 °C to 0,9 °C |  |
| 10 | Platinum resistance thermometers | Temperature, degrees Celsius (°С) | -45 °C to 150 °C | 0,05 °C | МК 7.02 01-10:2018 direct method |
|  |  |  | 150°C to 300 °C | 0,05 °C to 0,15 °C |  |
|  |  |  | 300 °C to 650 °C | 0,15 °C to 0,7 °C |  |
| 11* | Pressure gauges, vacuum gauges, pressure-vacuum gauges, calibrators, pressure transmitters with uniform output signal. | Pressure bar (bar) operation medium - gas | -0,95 bar to 1,5 bar | 1.10-3 bar | МК 7.02 01-11:2020 direct and comparative method |
|  |  |  | 0 bar to 35 bar | 5.10-3 bar to 15.10-3 bar |  |
|  |  | Pressure bar (bar) operation medium - fluid | 0 bar to 400 bar | 0,2 bar |  |

* At the calibration of converters with unified input or output signal under item 5 and item 11, electrical signals are used:

- for DCI in the range: 0 mA DC to 20 mA DC

- for DCU in the range: 0 V DC to 10 V DC

**The calibration of measuring instruments is carried out in the calibration laboratory premises or on-site at the customer.

References:

1.МК 7.02 01-01:2020 - Procedure for calibration of voltmeters and calibrators for DC and AC voltage.

2.МК 7.02 01-02:2020 - Procedure for calibration of ammeters and calibrators for DC and AC current

3.МК 7.02 01-03:2018 - Procedure for calibration of DC resistance standards, DC resistance calibrators and DC resistance decade boxes.

4.МК 7.02 01-04:2020 - Procedure for calibration of ohmmeters.

5.МК 7.02 01-05 :2020 - Procedure for calibration of secondary converters with DCU, ACU, DCI, ACI and resistance input, and with DCU, DCI and resistance output.

6.МК 7.02 01-06:2018 - Procedure for calibration of digital thermometers.

7.МК 7.02 01-07:2019 - Procedure for calibration of Relative humidity meters, sensor-transmitters for relative humidity in an air chamber.

8.МК 7.02 01-08:2018 - Procedure for calibration of analogue and liquid thermometers.

9.МК 7.02 01-09:2018 - Procedure for calibration of thermocouples.

10.МК 7.02 01-10:2018 - Procedure for calibration of Platinum resistance thermometers.

11.МК 7.02 01-11:2020 - Procedure for calibration of pressure measuring and generating devices (vacuum gauges, manometers, pressure transmitters, pressure calibrators)