
SCOPE 20 ЛК
Kalibrator LAB EOOD, Calibration laboratory

Management  address: 

1404 Sofia, Gotse Delchev Blvd., Building 37, Entrance 2, floor 8, apt.47

Laboratory addresses: 

Office 1: 1404 Sofia, Manastirski Livadi - East Housing Estate, Green Residence, Building 200, Entrance 1, Office P-111

Office 2: 1618 Sofia, Ovcha Kupel Housing Estate, 98 President Lincoln Blvd.
To perform calibrating of:
	Type of the scope: Fixed 

	№
	Measuring instrument
	Measured value,

unit of measure
	Measurement Range
	Calibration and measurement capability (CMC)
	Calibration method

	1
	2
	3
	4
	5
	6

	1
	Tanks*
	

	
	1.1. Horizontal tanks made of metal, mounted permanently and on movable platforms
	Volume, m3

	to 400 m3

	0.10 %
	М_02:2018

(strapping method)

	
	1.2. Vertical

cylindrical tanks
	Volume, m3

	to 20,000 m3

	0.10 %
	М_01:2018

(strapping method)



	
	1.3. Horizontal, vertical tanks and tanks of different geometric shapes
	Volume, m3
	to 75 m3
	0.18 %
	М_03:2018

(volumetric method)

	2
	Flowmeters for water 
and liquids other than water **
	Volume, m3 at set flow rate (from 3 m3/h to 9 m3/h)
	from 0.002 m3 to 0.30 m3
	0.10 %
	М_04:2020

(weighing method)

	
	
	Volume, m3 at set flow rate (9 m3/h to 75 m3/h)
	from 0.30 m3  to 2.50 m3
	0.20 %
	

	3
	Metal measures**
	Volume, l
	from 10 l to 
50 l
	from 0.07 % to 0.016 %
	М_05:2019

(weighing method)


* The calibration of measuring tools is performed at the client's premises.

** The calibration of measuring tools is performed at the premises of the calibration laboratory.

References:

M_01:2018: Vertical tank calibration by geometric method (БДС ISO 7507-1:2016  Petroleum and liquid petroleum products. Calibration of vertical cylindrical tanks. Part 1: Strapping method);
M_02:2018: Horizontal tank calibration by geometric method (ISO 12917-1:2017 Petroleum and liquid petroleum products. Calibration of horizontal cylindrical tanks. Part 1: Manual methods);
M_03:2019: Tank calibration by volumetric method (ISO 4269:2001 Petroleum and liquid petroleum products. Tank calibration by liquid measurement. Incremental method using volumetric meters);
M_04:2020: Calibration of flowmeters for water and liquids other than water by weighing method; 
M_05:2019: Calibration of metal volume measures by weighing method (EURAMET Calibration Guide № 19 Version 3.0 (09/2018)
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